| 1.0 | OBJECTIVE : | To lay down the procedure for handling of Laboratory Incident in the Quality Control Laboratory. |
| --- | --- | --- |
| 2.0 | RESPONSIBILITY : | QC Officer / QC Chemist / QC Head / QA Executive |
| 3.0 | ACCOUNTABILITY : | QA Head |
| 4.0 | PROCEDURE : |  |
| 4.1 | Laboratory Incident may be defined as any unplanned or unwanted or uncontrolled event that occurred in the laboratory in the form of non-compliance to the designed systems or procedures at any stage of testing, and storage of drug product due to system failure or equipment breakdown or manual error. |  |
| 4.2 | A laboratory incident is an event in the laboratory that occurs for two primary reasons either due to analyst error or instrument error. |  |
| 4.3 | The source of the lab incidents includes the following examples :- |  |
| 4.3.1 | General Incidents :- |  |
|  | - Mistake in the calculation and reporting, - Any contamination or error during the sample preparation, standard preparation and storage of samples, - Usage of un-calibrated instruments like analytical balance, pH meter, conductivity meter etc. - Deviation or error in weighing, dilutions, standards etc., - Wrong entry of method of parameter, sequence etc. in the software system, - Improper sample handling during analysis, - Glassware breakage having sample or standard, - Discontinue analysis for electrical problems or due to urgent re-planning of work, - Sample spillage during Loss on drying, Sulphated ash, moisture content etc., - Tested for parameter other than given in the specification, - Error in diluent volume or diluent use, - Wrong injection volume injected, - Entry missing in instrument logbook, - Mobile phase preparation error or improper mixing of mobile phase, - Improper baseline observed and improper integration observed before release of product / materials, - Incorrect instrument set up and run, - Wrong detector wavelength |  |
| 4.3.2 | Chromatography :- |  |
|  | - Injection carryover, - Additional / unknown impurities, - Shift in retention time / relative retention time, - Improper peak shape and peak splitting, - Baseline drift |  |
| 4.3.3 | Spectroscopy :- |  |
|  | - Extra peaks in the spectrum, - Low correlation with standard spectrum |  |
| 4.3.4 | Microbiology :- |  |
|  | - Wrong media preparation, - Borderline results, - Other unplanned unwanted error |  |
| 4.3.5 | Instrumental Error :- |  |
|  | - Column issue - Software malfunctioning, - Hardware error, - Leak detection, - Frit / cartridge choke, - Failure of lamp intensity, - Column leak, poor plate count, poor resolution, - Connection problem between software and instrument, - System failure observed during the analysis because of column high pressure |  |
| 4.4 | Incident Procedure :- |  |
| 4.4.1 | When an analyst obtains atypical results, or when the analyst becomes aware of laboratory errors that were committed during testing, he or she shall annotate the atypical results or known laboratory error in the appropriate analytical protocol / document and notify section in-charge immediately. |  |
| 4.4.2 | Analyst shall preserve all starting materials, solvents, intermediate and final solutions, and standards used in the testing. These materials will be retained and preserved until completion of the investigation or until it has been determined that they are no longer useful. |  |
| 4.4.3 | Analyst shall inform to QA personnel regarding observed incident. QA personnel shall enter the details of sr. no., date, incident report no., name of product / material, batch no. / AR no., description of incident, analyst, checked by and remark (if any) in the “Laboratory Incident Register” as per annexure-I. |  |
| 4.4.4 | The analyst shall obtain “Laboratory Incident Report” as per annexure-II from QA department and shall fill the relevant details into incident report. Based on review of incident, analyst shall intimate to QA for logging of incident report in the “Laboratory Incident Register” as per annexure-I. |  |
| 4.4.5 | Any laboratory incident observed shall be logged within one working day from the discovery of incident. |  |
| 4.4.6 | QA person shall assign a sequential number to the incident report as per below :- |  |
|  | LIR/YY/NNN |  |
|  | Where, |  |
|  | LIR represents Laboratory Incident Report, YY represents last two digit of respective year, i.e., 21 for year 2021, 22 for year 2022 etc., NNN represents continuous serial number for respective year starting from 001. |  |
| 4.4.7 | For example, first laboratory incident registered in year 2021 shall be numbered as LIR/21/001. |  |
| 4.4.8 | The QC Head/Designee will review the results, analytical protocol, starting materials and all solutions, raw data, and calculations used to generate the final results. |  |
| 4.4.9 | Laboratory Incident Report shall be completed together by the analyst, reviewer and analyst. The QC Head/Designee will interview the analyst, and will check relevant documents, in order to answer questions which will help to identify whether common or obvious laboratory errors were the root cause of the Incident/atypical result. |  |
| 4.4.10 | The investigation findings shall be documented in the incident Report. |  |
| 4.4.11 | If the original test solution is suspected, re-dilution of the stock or intermediate sample preparations can be re-analyzed. |  |
| 4.4.12 | If all re-analyzed, re-diluted results do not confirm the original result, and meets acceptance criteria, then dilution error may be the assignable cause. If this is found to be the case, then the re-dilution results from the re-analysis are accepted. The original result is invalidated and the investigation is closed. |  |
| 4.4.13 | If the re-dilution results do not confirm the atypical results, other factors such as insufficient mixing, sonication, shaking, etc., may be suspected as the assignable cause. The original atypical result shall be invalidated. Re-test the sample according to the method. |  |
| 4.4.14 | If any of the re-dilution results confirm the original atypical result, a full investigation is required to determine if the original results are valid or not. |  |
| 4.4.15 | Analyst shall attach all the relevant data with the incident report. |  |
| 4.4.16 | When an incident has been observed during analysis on HPLC, analyst shall fill the ‘HPLC Checklist for Laboratory Incident’ as per Annexure-III. |  |
| 4.4.17 | Reviewer shall propose the corrective actions in the incident report which shall be approved by the QC Head. |  |
| 4.4.18 | Head QC shall describe the details of preventive action taken in the report which shall be approved by QA Head. |  |
| 4.4.19 | Laboratory incident shall be closed within 30 working days from date of incident reported. The reason for any extension shall be approved by QA head and documented. |  |
| 5.0 | ANNEXURES:- |  |
| 5.1 | Annexure-I : Laboratory Incident Register |  |
| 5.2 | Annexure-II : Laboratory Incident Report |  |
| 5.3 | Annexure-III : HPLC Checklist for Laboratory Incident |  |
| 6.0 | REFERENCES:- NA |  |
| 7.0 | REVISION HISTORY:- |  |
| Sr. No.; Revision No.; Date of Revision; SOP No.; Supersedes No.; Reason for Revision. 02; 07/11/2022; QAGE113-02; QAGE113-1; - Change in Firm Name from “Indiana Ophthalmics” to “Indiana Ophthalmics LLP”. - Addition of Annexures, References and Revision History as Section 5.0, 6.0 and 7.0 respectively in the SOP.. |