| 1.0 | OBJECTIVE : | To lay down the procedure for receipt, performance and handling of High Performance Liquid Chromatography Column. |
| --- | --- | --- |
| 2.0 | ## RESPONSIBILITY : | QC Chemist, QC Head |
| 3.0 | ### ACCOUNTABILITY : | QA Head |
| 4.0 | PROCEDURE : |  |
| 4.1 | QC person shall receive the new column and shall verify the physical condition of column against indent or bill. |  |
| 4.2 | QC person shall condition the column as per defined procedure and shall check the performance of new column as per following steps after conditioning of column:- |  |
| 4.2.1 | Wash the Column 100% by using degas solvent such as HPLC grade Methanol / Acetonitrile for 30 min. |  |
| 4.2.2 | Wash the Column 100% with degas HPLC / filtered and degas distilled water up to 30 min. |  |
| 4.2.3 | Wash the Column with 60:40 ratio of degas Purified water and solvent up to 30 min. |  |
| 4.3 | Perform the System Suitability of the Column as per supplier COA for which Column intended and the observation shall be recorded as per Annexure-I (Column Performance Record). |  |
| 4.4 | If Column System Suitability criteria fails, immediately reject the Column and the same shall be sent to Purchase Department. |  |
| 4.5 | QC person shall assign unique Identification Number to the column if the performance of column found satisfactory. |  |
| 4.6 | Column shall be numbered as per following : ICH-XX Where, I stands for Indiana Ophthalmics LLP, C stands for Column, H Stands for HPLC, XX stands for continuous serial number starts with 01 & so on. |  |
| 4.7 | QC person shall take the entry of new column in “Column Issuance and Destruction Record” as per Annexure-II. QC person shall make an entry of product / material name, total injection, cumulative injection and column washing time of the particular column as per Annexure-III (HPLC Column Logbook). |  |
| 4.8 | Performance and Regeneration of in-use Column :- |  |
| 4.8.1 | Analyst shall re-wash the column with 100% degas water for 1 hr. & then the ratio of water & solvent (60:40 water: solvent) for 1 to 2 hour if performance or system suitability of column fails during routine analysis. |  |
| 4.8.2 | QC person shall check the performance of column after washing with the first injection of standard as per given method of analysis, should meet the system suitability criteria as given in method of analysis of the respective product. |  |
| 4.8.3 | Column to be regenerated if found poor performance after re-washing with Acetonitrile / Methanol / THF/ Butanol / IPA as per requirement & description of the column and then wash with 60:40 ratio of degas water: solvent for 1-2 hr. |  |
| 4.8.4 | QC Head shall retrieve and dispose failure old column with detail entry in issuance and destruction record and shall take in own custody to avoid reuse of column if it is not meeting the criteria of tailing factor, theoretical plates, resolution, fronting, splitting / damage etc. |  |
| 4.8.5 | After disposal of old column QC head shall issue the new column with given identification number of new column. |  |
| 4.9 | Cleaning and Maintenance of Column during Analysis :- |  |
| 4.9.1 | Analyst shall wash the in-use column with the ratio of 60:40 water: solvent after analysis for 30 to 60 min. (use the solvent whichever is used in the testing). |  |
| 4.9.2 | If a high concentration of buffer is used during analysis then wash the column by filtered 50°C hot water & then 100 ml with degas solvent. |  |
| 4.9.3 | Analyst shall ensure that column should wash with a mixture of methanol / acetonitrile or as per the specification recommended by the manufacturer if stored for a long time. |  |
| 4.10 | Precautions :- |  |
| 4.10.1 | Degassed solvents, water and mobile phase should be used. |  |
| 4.10.2 | Use clean filters. |  |
| 4.10.3 | Flush several volumes of 100% methanol (or similar) through the HPLC system after washing with water prior to shutting down. |  |
| 4.10.4 | Clean pump, Injector loop and syringe with hot water in a week. |  |
| 4.10.5 | Filter sample to be injected to avoid deposition of contamination at column end. |  |
| 4.10.6 | The column should be stored at a room temperature. |  |
| 4.10.7 | Protect the column from knocks and bumps. |  |
| 4.11 | Destruction of HPLC Column :- |  |
| 4.11.1 | Discard the column if it is not meeting the criteria of tailing factor, theoretical plates, resolution, fronting, splitting / damage etc. |  |
| 4.11.2 | QC person shall make an entry in the Column Issuance and Destruction record. |  |
| 5.0 | ANNEXURES:- |  |
| 5.1 | Annexure-I : Column Performance Record |  |
| 5.2 | Annexure-II : Column Issuance and Destruction Record |  |
| 5.3 | Annexure-III : HPLC Column Logbook |  |
| 6.0 | REFERENCES:- NA |  |
| 7.0 | REVISION HISTORY:- |  |
| Sr. No.; Revision No.; Date of Revision; SOP No.; Supersedes No.; Reason for Revision. 02; 04/10/2022; QAGE123-02; QAGE123-1; - Change in Firm Name from “Indiana Ophthalmics” to “Indiana Ophthalmics LLP”. - Addition of Annexures, References and Revision History as Section 5.0, 6.0 and 7.0 respectively in the SOP.. |