| 1.0 | ### OBJECTIVE : | To describe the process of Post-Market Surveillance (PMS). It defines roles & responsibilities as well as the interfaces needed to other economic operators, notified bodies and competent authorities. The purpose of this procedure is to provide for systematic data collection, analysis, reporting and review of information about the safety and efficacy of products that have been released to the market. |
| --- | --- | --- |
| 2.0 | ### SCOPE : | This procedure is applicable for post marketing surveillance and vigilance of medical devices or finished product manufactured at Indiana Ophthalmics LLP. |
| 3.0 | ## RESPONSIBILITY : | Marketing in-charge / QA Head / Management Representative |
| 4.0 | ACCOUNTABILITY : | QA Head |
| 5.0 | PROCEDURE : |  |
| 5.1 | This procedure covers all products that have been transferred from product development to post-market surveillance. This is the primary document meeting the applicable regulatory requirements for Post-Market Surveillance and a supporting document for Risk Management. |  |
| 5.2 | Post-market surveillance is a set of activities conducted by manufacturers, to collect and evaluate experience gained from products that have been placed on the market, and to identify the need to take any action. Post-market surveillance is a crucial tool to ensure that product continues to be safe and well performing and to ensure actions are undertaken if the risk of continued use of the product outweighs the benefit. |  |
| 5.3 | The terms post-market surveillance, vigilance and market surveillance are closely linked. Experiences gathered (feedback) on the use of medical devices/ products by users/ distributors are reported to manufacturers. Manufacturers report certain incidents to NRAs (National regulatory Authority)/ competent authority and keep them updated on the actions taken. The NRA will review the investigation undertaken by manufacturers and the further actions taken. This is connected to the market surveillance responsibilities of NRAs/ competent authority. |  |
| 5.4 | Post-market surveillance:- Systematic process to collect and analyze experience gained from products that have been placed on the market Source: |  |
| 5.4.1 | For each device, manufacturers need to plan, establish, document, implement, maintain and update a post-market surveillance system in a manner that is proportionate to the risk class and appropriate for the type of device. |  |
| 5.4.2 | The post-market surveillance system need to be suited to actively and systematically gathering, recording and analyzing relevant data on the quality, performance and safety of a device throughout its entire lifetime, and to drawing the necessary conclusions and to determining, implementing and monitoring any preventive and corrective actions. |  |
| 5.4.3 | Data gathered by the manufacturer's post-market surveillance system shall in particular be used: - to update the benefit-risk determination and to improve the risk management, - to update the design and manufacturing information, the instructions for use and the labeling, - to update the clinical evaluation, - to update the summary of safety and clinical performance, - for the identification of needs for preventive, corrective or field safety corrective action, - for the identification of options to improve the usability, performance and safety of the device, - when relevant, to contribute to the post-market surveillance of other devices, - to detect and report trends |  |
| 5.4.4 | If, in the course of the post-market surveillance, a need for preventive or corrective action or both is identified, the manufacturer need to implement the appropriate measures and inform the competent authorities concerned and, where applicable, the notified body. Where a serious incident is identified or a field safety corrective action is implemented, it has to be reported. |  |
| 5.5 | PMS document pertains to the objectives and processes for post-market surveillance for products conducted by manufacturers with the assistance of their economic operators, as well as market surveillance conducted by regulators, and the role of other stakeholders in these processes. It describes the measures taken to ensure the ongoing compliance of products with the requirements for safety, quality and performance after they are placed on the market. |  |
| 5.6 | Post market Surveillance (PMS) is the requirement that manufacturer have to monitor their products after they have been cleared for sale and are in use by members of the public. |  |
| 5.7 | A basic principle of post-market surveillance is Pre-market evaluation of product quality, safety and performance is conducted by manufacturers of products prior to entry on to the market. |  |
| 5.8 | This procedure identifies and implements a system for gathering and analyzing post-production feedback on products to ensure customer requirements are met. |  |
| 5.9 | Roles and Responsibilities:- |  |
| 5.9.1 | Manufacturer’s responsibilities: Although products are designed, developed, manufactured and distributed on the global market after thorough pre-market evaluation, residual risks regarding safety and performance will remain throughout the product’s lifetime. Design and development activities for products ensure the residual risks are acceptable with respect to anticipated benefits before the product is released on to the market. It remains important to continue to collect and evaluate information on the product during production and post-production to meet requirements for the monitoring of products and processes and to ensure the residual risks remain acceptable with respect to benefits. Appropriate processes allow for early detection of any undesirable effects. |  |
| 5.9.2 | Quality and regulatory team is responsible for composing post-market surveillance reports, or re-evaluation of existing risk control measures, and presenting information to top management. |  |
| 5.9.3 | Top management :- Top management has the overall responsibility for providing an adequate organizational structure including sufficient qualified resources. Top management oversees the effectiveness of the PMS system and has discretion to initiate corrections, corrective and preventive actions as needed as per SOP No. QAGE115 (Management review). |  |
| 5.9.4 | Sales & Distribution :- The sale department should supply the sales information and customer investigation information to quality department. |  |
| 5.9.5 | The quality department should collect the customer complaints information and other feedback and maintain all documents. |  |
| 5.9.6 | Stakeholder’s roles and responsibilities: Post-market surveillance should be implemented by every manufacturer, at least in its most basic form, as a system to monitor, collect, evaluate and react to feedback. Other economic operators, such as importers, distributors and authorized representatives, play an important supportive role to ensure feedback from users reaches the manufacturer, including overcoming language barriers. |  |
| 5.9.7 | It is responsibility of respective distributors/ parties to appoint authorized person for pharmacovigilance activities at importing country. |  |
| 5.9.8 | Sponsors/ Distributors must inform the regulatory authority of respective country, the name and contact details of the person (‘the nominated contact person’) responsible for fulfilling the sponsor’s reporting requirements for the respective country. |  |
| 5.9.9 | Qualified personnel for pharmacovigilance is responsible for conduction scientific research relevant to pharmaceutical within their territory, for composing post-marketing surveillance reports, customer investigation information to quality department. |  |
| 5.9.10 | Qualified personnel for pharmacovigilance for distributors/ parties provide necessary information, data, feedback, complaints & other details to respective manufacturer. |  |
| 5.10 | Manufacturers and their economic operators:- |  |
| 5.10.1 | Implement a system for post-market surveillance: An effective post-market surveillance system should include both active and passive collection of post-market information. Collecting and evaluating feedback are critical. |  |
| 5.10.2 | Establish if reporting to NRA is required: The manufacturer needs to establish if reporting to the NRA is required. Initial, follow-up and final investigation reports should contain all details of any investigation conducted. |  |
| 5.10.3 | If required, undertake root cause analysis: The manufacturer should perform root cause analysis to establish the root cause for the issue, allowing adequate action to be initiated |  |
| 5.10.4 | If required, make a correction: Corrections and corrective/preventive actions may also be required to protect public safety. |  |
| 5.10.5 | If required, implement corrective and/or preventive actions. |  |
| 5.11 | The most basic form of post-market surveillance, which shall always be performed, is reactive post-market surveillance. Reactive post-market surveillance is done through collection and evaluation of feedback from distributors & customers. |  |
| 5.12 | Relevant procedures such as customer feedback procedure (SOP No.: PEGE24), market complaints (SOP No.: QAGE02), recall (SOP No.: QAGE03) and advisory notice (SOP No.: QAGE128), self-inspection (internal audit) (SOP No.: QAGE06), corrective and preventive action (SOP No.: QAGE124), etc. are available to strengthen the process of post-marketing surveillance. |  |
| 5.13 | Collect Feedback:- |  |
| 5.13.1 | As one of the measurements of the effectiveness of the quality management system, the company has gathered and monitors information relating to whether the company has met customer requirements. |  |
| 5.13.2 | Receiving and acting upon user or other feedback is the most basic form of post-market surveillance that must always be performed by the manufacturer, irrespective of their resources. |  |
| 5.13.3 | The methods for obtaining and using this information’s shall be documented as a form of customer feedback, distributor feedback, market complaint register etc. |  |
| 5.13.4 | Feedback received from the customers or distributors shall be handled as per SOP No. PEGE24 (Customer Feedback Procedure). |  |
| 5.13.5 | Feedback system helps in continuous improvement of the organization by identifying and implementing any changes necessary to ensure and maintain the continued suitability and effectiveness of the quality management system through the use of the quality policy, quality objectives, audit results, analysis of data, corrective and preventive actions and management review. |  |
| 5.14 | Complaint Handling:- |  |
| 5.14.1 | Market complaint received shall be handled as per SOP No. QAGE02 (Handling of Market Complaints). |  |
| 5.14.2 | These procedures shall include below requirements and responsibilities: |  |
|  | - Receiving and recording information, - Investigation complaints, - Determining the need to report the information to the appropriate regulatory authorities, - Handling of complaint related product, - Determining the need to initiate corrective & preventive actions. |  |
| 5.14.3 | Market complaint handling procedure is intended to investigate customer complaints to determine the cause and take the necessary remedial action, determine methods to preclude re-occurrence of the problem. |  |
| 5.14.4 | If any complaint is not investigated, justification shall be documented. |  |
| 5.14.5 | Complaint handling records shall be maintained according SOP No. QAGE17 (Document Control Procedure). |  |
| 5.15 | Reporting to Regulatory Authorities:- |  |
| 5.15.1 | If regulatory requirements require notification of complaints that meet specified reporting criteria of adverse events or issuance of advisory notices, the company shall inform or notify regulatory authorities. |  |
| 5.15.2 | If any malfunction or deterioration in the characteristics and/or performance of a product, as well as any inadequacy in the labeling or the instructions for use which, directly or indirectly, might lead to or might have led to the death of a patient, or user or of other persons or to a serious deterioration in their state of health, then it shall be covered into recall and advisory notice procedures. |  |
| 5.16 | Internal Audit:- |  |
| 5.16.1 | Company shall conduct internal audits at planned intervals (every six months) to determine whether the quality management system: - Conforms to planned and documented arrangements, requirements of this international standards, quality management system requirements established by the company and applicable regulatory requirements, - Is effectively implemented & maintained. |  |
| 5.16.2 | Company has defined a brief procedure for internal audit as per SOP No.: QAGE06 “Internal Facility Audit (Self Inspection)”. |  |
| 5.17 | CAPA procedure is intended to ensure that corrective and preventive action is a vital closed loop system of root cause analysis, documented action, verification of effectiveness, and prevention of recurrence. |  |
| 5.18 | In post marketing surveillance, product which does not conform to product requirements shall be identified and controlled to prevent its unintended use or delivery. |  |
| 5.19 | Effective arrangements for communicating with customers shall be determined and implemented in relation to:- - Product information, - Enquiries, contracts or order, - Customer feedback, including customer complaints |  |
| 5.20 | PMS Plan:- |  |
| 5.20.1 | Typically the PMS Plan for each product is initiated after the approval of design outputs (i.e., "design freeze") and the PMS Plan is approved at the end of the design transfer process along with commercial release of the product. |  |
| 5.20.2 | Individual PMS documents are controlled according SOP No. QAGE17 (Document Control Procedure). |  |
| 5.20.3 | The following elements should be considered in the preparation of the PMS Plan for a device or device family:- |  |
|  | - Information from research and development activities, within Indiana Ophthalmics LLP or at any contract site, - Information from users of the products, including customer complaints and any other means by which feedback may be sought, e.g. customer feedback, - Frequency and nature of clinical evaluation, e.g. post-market clinical studies or published literature, - Information relating to market experience of similar products or product with similar intended use, during cleaning evaluation report. - Information on new or amended standards, regulations, regulatory guidance. |  |
| 5.21 | Other Sources of post-market information:- |  |
| 5.21.1 | Information about the product may become available in sources of post-market information other than user feedback, for example, through literature. |  |
| 5.21.2 | The manufacturer shall choose the appropriate sources in line with the objectives of the post-market surveillance plan. A non-exhaustive overview of possible data sources includes: |  |
|  | - Incidents reported - Preventive maintenance & repair - Return & recall - Post market clinical follow up (PMCF) or justification as to why a PMCF is not applicable is mentioned in CER. - Advisory notice - Literatures - Internal audits |  |
| 5.22 | Post-market surveillance plan shall be prepared. PMS plan shall include the following parameters :- |  |
|  | - Information Gathering (Data Collection) - Handling Measures (Data Analysis) - Review Requirements - Vigilance Reporting |  |
| 5.23 | Information Gathering (Data Collection):- |  |
| 5.23.1 | This shall include set-up of telephone numbers to collect customer complaints, record, analyze and deal with them in time. (Telephone numbers available on website). |  |
| 5.23.2 | Sales department shall issue the market information feedback record form to all customers, when the customer meets some abnormal situation, they shall fill in the form and return to the company in time. |  |
| 5.23.3 | Sampling survey by telephone in order to know whether there are any doubt and improvement suggestions about the use of product and to get knowledge about the satisfaction of customers for product, make a call to some customers, shall be carried by the marketing team. |  |
| 5.23.4 | Reactive post-market surveillance, based on collecting feedback, should always be in place. The manufacturer shall choose the appropriate data sources to allow the fulfillment the objectives of the post-market surveillance plan. For example, to ensure that the product remains state-of the-art, actively collecting data on similar product and procedures from literature is required. |  |
| 5.23.5 | On receipt of the finished goods, by a representative of the party, the consignment shall be inspected for any visible damage from outside and shall record the observation. All the relevant entries shall be recorded as per Annexure-VII (Finished Goods Receipt Checklist). |  |
| 5.24 | Handling Measures (Data Analysis):- |  |
| 5.24.1 | All the details / information gathered and collected for the product during information gathering shall be deal actively by providing corrective and preventive measures. |  |
|  | Complaints are collected, classified and analyzed by Quality Assurance Department as per SOP No. QAGE02. Complaints related to product quality are forwarded to Quality Assurance for evaluation and investigation of root causes. - Monitoring of adverse event and control procedure of advisory notification - Control procedures of corrective and preventive actions - Product recall management system - Market research and analysis management system - Quality complaints management - Product returns management regulations |  |
| 5.25 | Review Requirements:- |  |
| 5.25.1 | The collected information shall be reviewed, and the root causes shall be investigated. The data shall be analysed for trends and for assurance that no unacceptable or previously unrecognized hazards occurred. |  |
| 5.25.2 | Any deficiencies identified in the current product shall be feedback to the relevant design team and discussed with senior management to determine appropriate corrective/preventive actions. |  |
| 5.25.3 | PMS data collection and review is conducted as on-going activity. The literature review and CER update shall be performed every three year. |  |
| 5.25.4 | PMS information shall be discussed with senior management during Management Review Meeting as per SOP No. QAGE115. |  |
| 5.26 | Vigilance Reporting:- |  |
| 5.26.1 | It is responsibility of respective distributors/ parties to appoint authorized person for pharmacovigilance activities at importing country. |  |
| 5.26.2 | Qualified personnel for pharmacovigilance is responsible for conduction scientific research relevant to pharmaceutical within their territory, for composing post-marketing surveillance reports, customer investigation information to quality department. |  |
| 5.26.3 | Qualified personnel for pharmacovigilance for distributors/ parties provide necessary information, data, feedback, complaints & other details to respective manufacturer. |  |
| 5.26.4 | Company has also appointed responsible person for pharmacovigilance activities from regulatory affairs department. |  |
| 5.26.5 | Pharmacovigilance person from regulatory department is responsible for compilation of all data, feedback & information. |  |
| 5.26.6 | Any complaint related to product shall be forward to quality assurance department. |  |
| 5.26.7 | Incidents, adverse events feedback and warning shall be reported to Competent Authorities and Notified Body. |  |
| 5.27 | Post-Market Surveillance plan of the product shall be prepared as per annexure-I (Post-Marketing Surveillance Plan). |  |
| 5.28 | Periodic safety update report (PSUR):- |  |
| 5.28.1 | Manufacturers need to prepare a periodic safety update report (‘PSUR’) for each device and where relevant for each category or group of devices summarizing the results and conclusions of the analyses of the post-market surveillance data gathered as a result of the post-market surveillance plan & description of any preventive and corrective actions taken. Throughout the lifetime of the device concerned, that PSUR shall set out: |  |
| 5.28.2 | (a) the conclusions of the benefit-risk determination; (b) the main findings of the PMCF; and (c) the volume of sales of the device and an estimate evaluation of the size and other characteristics of the population using the device and, where practicable, the usage frequency of the device. (Data from EU countries, within origin of countries & third countries) |  |
| 5.28.3 | PSUR shall be updated when necessary and at least every two years. PSUR is an output of PMS activities. |  |
| 5.28.4 | Benefits of PSUR: - Provide an evaluation of the risk-benefit ratio, - Constant update of safety reports, - Provide insight into the worldwide safety data, - Risk identifications, - Corrective and preventive actions. |  |
| 5.28.5 | Periodic Safety Update Report (PSUR) of the product shall be prepared as per annexure-VI (Periodic Safety Update Report (PSUR)) which includes below points: |  |
| 5.28.5.1 | Introduction: Provide introduction of product / device which includes product name, international unpatented name, ATC code, Dosage Code, Product Description, Intended Use, Composition, Dose & Administration, Side Effects, Precautions/ warning etc. |  |
| 5.28.5.2 | Executive Summary: Highlight which issues around safety and performance of concerned device which are identified under section executive summary. |  |
| 5.28.5.3 | Updated data on actions taken by regulatory authority or owner of registration in the purpose of safety provision for device application under this section. |  |
| 5.28.5.4 | Impact on clinical evaluation: benefit and performance: provide details like Are the claims made for expected benefits and performance being met? ( ) Yes ( ) No If not, reasons for deviation shall be provide under this section. |  |
| 5.28.5.5 | Impact on risk management: risks and benefits determination: provide details like The medical device is safe and the risks correspond to the assumptions in the risk management file: ( ) Yes ( ) No. If not, describe new risks that have been identified: |  |
| 5.28.5.6 | Results Of Post-Market Surveillance Activities: |  |
|  | Summary of Incidents and Problems: Describe incidents that happened with device and actions taken (e.g. adverse event reports to authorities, field safety notices to customers, field safety actions). |  |
|  | Observed undesirable side effects: Describe incidents that happened with device and actions taken. |  |
|  | List of CAPAs: The following CAPAs resulted from post-market surveillance activities during product lifetime and the following CAPAs were initiated during this surveillance period: |  |
|  | Summary of Customer Complaints and Problems: |  |
|  | Summarize the number of customer complaints, predominant subjects of feedback and whether there was anything of importance regarding safety. |  |
|  | Summary of Post-Market Clinical Follow-Up Findings: Describe the main findings that were derived during post-market clinical follow-up. |  |
|  | Summary of sales data: the volume of sales of the device and an estimate evaluation of the size and other characteristics of the population using the device and, where practicable, the usage frequency of the device. (Data from EU countries, within origin of countries & third countries) |  |
| 5.28.5.7 | # Trend Analysis: Describe trends that identified according during post-market surveillance plan. |  |
| 5.29 | Procedure for reporting changes to Notified Body and Competent Authorities:- |  |
| 5.29.1 | The exchange of information between Indiana Ophthalmics LLP and notified bodies and competent authorities/ NRAs include, |  |
|  | - Reporting of serious incidents, trend reporting and field safety corrective actions and notices |  |
| 5.29.2 | Reporting of any change to the notified body and competent authorities shall be done through the “Change Notification Form”. This form shall be downloaded from the website, https://htcert.com/, under section:- CertificationNotification of ChangesSubmit a Substantial Change Notification. |  |
| 5.30 | Reporting of serious incidents and field safety corrective actions:- |  |
| 5.30.1 | Manufacturers of devices made available on the Union market, other than investigational devices, shall report, to the relevant competent authorities, the following: |  |
| 5.30.1.1 | (a) any serious incident involving devices made available on the Union market, except expected side-effects which are clearly documented in the product information and quantified in the technical documentation and are subject to trend reporting; (b) any field safety corrective action in respect of devices made available on the Union market, including any field safety corrective action undertaken in a third country in relation to a device which is also legally made available on the Union market, if the reason for the field safety corrective action is not limited to the device made available in the third country. |  |
| 5.31 | Timescale for the Initial Reporting of an Incident:- |  |
| 5.31.1 | Upon becoming aware that an event has occurred and that one of its device may have caused or contributed to that event, the Manufacturer must determine whether it is an Incident. |  |
| 5.31.2 | The following time lines shall apply in a case of the severity of the serious incident: |  |
| 5.31.2.1 | Manufacturer need to report any serious incident as referred to in point (a) immediately (without any delay that could not be justified) after the manufacturer have established causal relationship between that incident and their device or that such causal relationship is reasonably possible and not later than 15 days after they become aware of the incident to the relevant national regulatory authorities / national competent authority. |  |
| 5.31.2.2 | Event of a Serious public health threat: The Report need to provide immediately (without any delay that could not be justified), and not later than 2 days after the manufacturer becomes aware of that threat to the relevant national regulatory authorities / national competent authority. |  |
| 5.31.2.3 | Event of Death or Unanticipated serious deterioration in a person’s state of health: The Report need to provide immediately (without any delay that could not be justified) after the manufacturer has established or as soon as it suspects a causal relationship between the device and the serious incident but not later than 10 days after the date on which the manufacturer becomes aware of the serious incident to the relevant national regulatory authorities / national competent authority. |  |
| 5.31.3 | Where necessary to ensure timely reporting, the manufacturer need to submit an initial report that is incomplete followed up by a complete report. |  |
| 5.31.4 | If, after becoming aware of a potentially reportable incident, the manufacturer is uncertain about whether the incident is reportable, it shall nevertheless submit a report within the timeframe required in accordance with above timelines. |  |
| 5.31.5 | Except in cases of urgency in which the manufacturer needs to undertake field safety corrective action immediately, the manufacturer shall, without undue delay, report the field safety corrective action referred to in point (b) in advance of the field safety corrective action being undertaken. |  |
| 5.31.6 | For similar serious incidents that occur with the same device or device type and for which the root cause has been identified or a field safety corrective action implemented or where the incidents are common and well documented, the manufacturer need to provide periodic summary reports instead of individual serious incident reports. |  |
| 5.31.7 | The manufacturer should also inform to notified body & authorized representative of all countries affected at the same time. |  |
| 5.32 | A template for manufacturer’s incident reports as per annexure-II (Manufacturer’s Incident Report Form). |  |
| 5.33 | Field Safety Corrective Action (FSCA):- |  |
| 5.33.1 | An action taken by a manufacturer to reduce the risk of death or serious deterioration in the state of health associated with the use of products. Such actions should be notified via a field safety notice (FSN). |  |
| 5.33.2 | The Medical Device Directives require the Manufacturer to report to the National Competent Authority any technical or medical reason leading to a systematic recall of devices/ products of the same type by the Manufacturer. Those reasons are any malfunction or deterioration in the characteristics and/or performance of a product, as well as any inadequacy in the instructions for use which might lead to or might have led to the death of a patient or USER or to a serious deterioration in his state of health. |  |
| 5.33.3 | FSCAs may include: - return of a products/ devices to the manufacturer or its representative; - Product/device modification; - Product/device exchange; - Product/ device destruction; - advice given by the manufacturer regarding the use of the product/device |  |
| 5.33.4 | An FSCA report should be submitted to NRAs of all the countries affected through an FSCA report, if required by national legislation. |  |
| 5.33.5 | The FSCA final report should contain the following information: - A final assessment of the root cause of the problem and proposed corrective action to reduce the chance of recurrence, e.g. redesign, update in the field, improved IFU etc. and the progress of the implementation of such actions. - The outcome of the reconciliation of the FSCA. |  |
| 5.33.6 | Field safety corrective action is a synonym for recall or withdrawal. |  |
| 5.33.7 | Company has well defined procedure for Recall and Withdrawn of the marketed products. |  |
| 5.33.8 | It is the responsibility of manufacturer to produce the product of consistent quality required for the intended use. |  |
| 5.33.9 | If the product is not found stable during its stability studies, it is the responsibility of manufacturer to withdraw the marketed formulation voluntarily. |  |
| 5.33.10 | A recall is a voluntary action taken by the manufacturer at any time to remove the defective drug product/ device from the market. Possible reasons for the product recall is listed below:- |  |
|  | a) If the product is found to be hazardous for human use. b) Due to loss of potency. c) Due to microbial contamination. d) Due to degradation of the product and when such product is found hazardous for human use. e) Any other reason found by regulatory authority. |  |
| 5.33.11 | Recall / withdrawal procedure is well described in SOP No.: QAGE03 “Guidelines for Recall & Withdrawal Procedure for Marketed Product”. |  |
| 5.33.12 | A template for manufacturer’s Field Safety Corrective Action reports as per annexure-III (Manufacturer’s Field Safety Corrective Action Report). |  |
| 5.34 | Field safety notice (FSN):- A communication sent out by a manufacturer or its representative to the product users in relation to FSCA. Note: FSN can also be non-safety related, e.g. quality-related, customer product information. |  |
| 5.34.1 | Field safety notices are an important means of communicating FSCA and safety information to users. |  |
| 5.34.2 | Manufacturers should inform affected users of any FSCA via FSN, and inform the relevant NRAs. Affected users will usually receive the FSN via their procurement agents or through economic operators, who must inform all users within their region of supply. |  |
| 5.34.3 | The FSN should include the following items - A clear title such as “Urgent Field Safety Notice” on the notice itself, and the subject line, if sent by e-mail. - The intended audience: clear statement about the intended recipient of the notice. - Concise description of product, product code, lot/ batch number(s). - A factual statement explaining the reasons for the FSCA, including a description of the problem. - A clear description of the hazards associated with the specific failure of the product and, where appropriate, the likelihood of occurrence, being mindful of the intended audience. - The recommended action(s) to be taken by the recipient of the FSN including any action(s) recommended for people that have previously used or been treated by the affected product, including recalls. - Where appropriate, include timeframes by which the action(s) should be taken by the manufacturer and user. - Designated contact point for the recipient of the FSN to obtain further information. |  |
|  | - The FSN should not include any: - Comments and descriptions that downplay the level of risk. - Information that is intended to promote a manufacturer or their product’s market visibility for the purposes of sales and marketing. |  |
| 5.34.4 | A template for field safety notice as per annexure-IV (Template for Field Safety Notice). |  |
| 5.35 | Trend Reporting:- |  |
| 5.35.1 | Manufacturers need to report, any statistically significant increase in the frequency or severity of incidents that are not serious incidents or that are expected undesirable side- effects that could have a significant impact on the benefit-risk analysis and which have led or may lead to risks to the health or safety of patients, users or other persons that are unacceptable when weighed against the intended benefits. |  |
| 5.35.2 | The significant increase need to be established in comparison to the foreseeable frequency or severity of such incidents in respect of the device, or category or group of devices, in question during a specific period as specified in the technical documentation and product information. |  |
| 5.35.3 | All feedback should be initially evaluated to determine if immediate action is required to protect public health and safety. |  |
| 5.35.4 | On identifying a significant increase or trend of events or Incidents should be made to the relevant National Competent Authority. |  |
| 5.35.5 | Unless otherwise required by national legislation, the manufacturer (or their economic operator) should send an investigation report to the NRA, in the following circumstances: |  |
|  | - Discovery of a serious public health threat. Serious public health threat: Any event type or device or product deficiency which could result in imminent risk of death, serious deterioration in the state of health, serious injury, or serious illness of more than one patient, user or other person that requires prompt remedial action. |  |
|  | - When use of a device led to: - – death of a user, patient/client or other person; - – serious deterioration in health of a user, patient/client or other person. |  |
|  | - No death or serious deterioration in health of a user, patient/client or other person occurred but might have. |  |
| 5.35.6 | This may include indirect harm, such as misdiagnosis, delayed diagnosis, delayed treatment, inappropriate treatment, absence of treatment or transfusion of inappropriate materials. |  |
| 5.35.7 | A template for manufacturer’s Trend reports Form as per annexure-V (Manufacturer’s Trend Report Form). |  |
| 5.36 | Analysis of serious incidents and field safety corrective actions:- |  |
| 5.36.1 | Following the reporting of a serious incident, the manufacturer need, without delay, perform the necessary investigations in relation to the serious incident and the devices concerned. This shall include a risk assessment of the incident and field safety corrective action. |  |
| 5.36.2 | Upon request by the national competent authority, manufacturers need to provide all documents necessary for the risk assessment. |  |
| 5.36.3 | The competent authority shall monitor the manufacturer's investigation of a serious incident. Where necessary, a competent authority may intervene in a manufacturer's investigation or initiate an independent investigation. |  |
| 5.36.4 | The manufacturer needs to provide a final report to the competent authority setting out its findings from the investigation. The report shall set out conclusions and where relevant indicate corrective actions to be taken. |  |
| 5.36.5 | The manufacturer need to ensure that information about the field safety corrective action taken is brought without delay to the attention of users of the device in question by means of a field safety notice. The field safety notice shall be edited in an official Union language. |  |
| 6.0 | ANNEXURES:- |  |
| 6.1 | Annexure-I : Post Marketing Surveillance and Vigilance Plan |  |
| 6.2 | Annexure-II : Manufacturer’s Incident Report Form |  |
| 6.3 | Annexure-III : Manufacturer’s Field Safety Corrective Action Report |  |
| 6.4 | Annexure-IV : Template for a Field Safety Notice (FSN) |  |
| 6.5 | Annexure-V : Manufacturer’s Trend Report Form |  |
| 6.6 | Annexure-VI : Periodic Safety Update Report (PSUR) |  |
| 6.7 | Annexure-VII : Finished Goods Receipt Checklist |  |
| 7.0 | REFERENCES:- |  |
| 7.1 | EN ISO 14971: 2019 Medical Devices - Application of risk management to medical devices |  |
| 7.2 | ISO TIR 24971:2013 Medical Devices - Guidance on the application of ISO 14971 |  |
| 7.3 | MEDDEV 2.7/1 Rev.3 2009 “Guidelines on medical devices Clinical Evaluation: A guide for manufacturers and Notified Bodies” |  |
| 7.4 | MEDDEV 2.12/2 Rev.2 January 2012 “Guidelines on medical devices Post-Market Clinical Follow-up Studies” |  |
| 7.5 | MEDDEV 2.12/1 Rev.8 January 2013 “Guidelines on medical devices Vigilance System” |  |
| 7.6 | Additional Guidance Regarding the Vigilance System as outlined in MEDDEV 2.12/1 Rev.8 July 2019 |  |
| 7.7 | REGULATION (EU) 2017/745 Of The European Parliament And Of The Council of 5 April 2017 on medical devices, amending Directive 2001/83/EC, Regulation (EC) No 178/2002 and Regulation (EC) No 1223/2009 and repealing Council Directives 90/385/EEC and 93/42/EEC |  |
| 8.0 | REVISION HISTORY:- |  |
| Sr. No.; Revision No.; Date of Revision; SOP No.; Supersedes No.; Reason for Revision. 03; 29/12/2022; QAGE127-03; QAGE127-2; - Change in Firm Name from “Indiana Ophthalmics” to “Indiana Ophthalmics LLP”. - Addition of Scope, Annexures, References and Revision History as Section 2.0, 6.0, 7.0 and 8.0 respectively in the SOP.. |