| 1.0 | ### OBJECTIVE : | To describe step-wise procedure for validation of Lab Solutions. |
| --- | --- | --- |
| 2.0 | ### SCOPE : | This procedure is applicable for validation of Lab solutions software installed in computers of Quality control laboratory at Indiana Ophthalmics LLP. |
| 3.0 | ## RESPONSIBILITY : | QC Chemist / QC Head |
| 4.0 | ACCOUNTABILITY : | QA Head |
| 5.0 | PROCEDURE : |  |
| 5.1 | This procedure describes guideline for the validation of Lab Solutions Software used to run the HPLC system in quality control department. |  |
| 5.2 | Validation of the Lab Solutions Software shall be carried out by External Validation Agency / validation consulting team. |  |
| 5.3 | The written agreement between Indiana Ophthalmics LLP and External Validation Agency shall be available. |  |
| 5.4 | The following documents shall be provided by the External Validation Agency as part of the validation of Lab Solutions Software to meet the necessary regulatory requirements as per GLP environment :- |  |
|  | - Validation Plan - GxP Assessment - Installation Qualification - Operational Qualification - Performance Qualification - Risk Assessment |  |
| 5.4.1 | Validation Plan :- |  |
| 5.4.1.1 | Validation plan shall be provided by the External Validation Agency which shall include, but not limited to, validation approach in details, roles and responsibility during execution of installation, operational and performance qualification. |  |
| 5.4.2 | ## GxP Assessment:- |  |
| 5.4.2.1 | This document shall provide impact assessment of the Lab Solutions software with respect to GxP requirements. |  |
| 5.4.3 | Risk Assessment:- |  |
| 5.4.3.1 | Risk Assessment document shall include the identification, managing and mitigation of the risks associated with the Lab Solutions software. |  |
| 5.4.3.2 | Two type of risk shall be considered while validating the Lab Solutions software. |  |
|  | - Initial Risk Assessment (IRA) - Functional Risk Assessment (FRA) |  |
| 5.4.3.3 | IRA (Initial Risk Assessment) shall identify overall risk of the system. It shall be prepared by considering all the functionalities that could affect factors such as product quality, safety and data integrity. |  |
| 5.4.3.4 | FRA (Functional Risk Assessment) shall be performed to identify potential risks. FRA shall be performed to identify each risk scenario for potential malfunction, impact on product quality, patient safety and data integrity harms based on severity, probability and detectability. |  |
| 5.4.4 | Installation Qualification (IQ) :- |  |
| 5.4.4.1 | Installation Qualification shall provide documented verification of key aspects for software and hardware installation. |  |
| 5.4.4.2 | Installation Qualification shall ensure that installation for the formal verification environment is setup and functioning as intended. |  |
| 5.4.4.3 | Installation Qualification shall be executed and documented for quality and production environment to demonstrate that components of the system are installed according to specifications. |  |
| 5.4.4.4 | The instruments used for the Installation Qualification shall have calibration certificate traceable to NIST/NPL or equivalent standards and shall have valid calibration on the date of protocol execution. |  |
| 5.4.4.5 | Installation Qualification shall include pre-approval of the protocol from representative of Indiana Ophthalmics LLP prior to executing. |  |
| 5.4.4.6 | Installation Qualification shall include, but not limited to, the following verification test :- |  |
|  | - Verification of Workstation Identification - Verification of Workstation Documentation - Verification of Workstation Hardware and Interface Components - Verification of Software Installed at Workstation - Verification of systems connection to the server - Verification of Environmental Conditions - Verification of Utilities Requirements |  |
| 5.4.4.7 | Discrepancies observed during the execution of Installation Qualification shall be documented and Corrective & Preventive actions shall be provided. |  |
| 5.4.4.8 | On completion of the Installation Qualification activity, post approval from the representative of Indiana Ophthalmics LLP is required to ensure effective installation as per user requirement. |  |
| 5.4.5 | Operational Qualification (OQ) :- |  |
| 5.4.5.1 | Operational Qualification shall be carried out to provide documented evidence that the system operates as required. Functional operation shall be verified to ensure that system operates as per requirement. |  |
| 5.4.5.2 | The result of the Operational Qualification ensures that the software behaves correctly and consistently in its operating environment. |  |
| 5.4.5.3 | Operational Qualification shall include pre-approval of the protocol from representative of Indiana Ophthalmics LLP prior to executing. |  |
| 5.4.5.4 | Operational Qualification shall include, but not limited to, the following verification test :- |  |
|  | - Verification of Start-up and Shutdown Procedure of Workstation - Verification of Operating System Security - Verification of Password Requirement for Application Software - Verification of User ID Requirement - Verification of User Role/ Authorization Control Requirement - Verification of Communication Failure Condition - Verification of Power Failure Condition - Verification of Standard Operating Procedure - Verification of User Training Record |  |
| 5.4.5.5 | Discrepancies observed during the execution of Operational Qualification shall be documented and Corrective & Preventive actions shall be provided. |  |
| 5.4.5.6 | On completion of the Operational Qualification activity, post approval from the representative of Indiana Ophthalmics LLP is required to ensure that system operates as per requirement. |  |
| 5.4.6 | Performance Qualification (PQ) :- |  |
| 5.4.6.1 | Performance Qualification (PQ) shall challenge and shall demonstrate system’s ability to perform functionality and operation of the system. |  |
| 5.4.6.2 | The Performance Qualification shall provide sufficient documentary evidence to demonstrate functional operation as specified. |  |
| 5.4.6.3 | Performance Qualification (PQ) shall be prepared providing documented verification that the equipment or system performance is as per requirement. |  |
| 5.4.6.4 | Performance Qualification shall include, but not limited to, the following verification test :- |  |
|  | - Verification of Operational Requirement - Verification of Functional Requirement - Verification of Report Output Requirement - Verification of Audit Trail - Verification of Data Export and Import - Verification of Data Backup and Restoration |  |
| 5.4.6.5 | Any discrepancies or deviation observed during the qualification activity of Lab Solutions software shall be documented in the respective qualification documentation. |  |
| 5.4.6.6 | The resolution of these items and/or any item outstanding that shall require further effort to resolve shall be included in the respective qualification documentation. |  |
| 5.4.6.7 | On completion of the Performance Qualification activity, post approval from the representative of Indiana Ophthalmics LLP is required to ensure that system operates as per requirement. |  |
| 5.5 | Acceptance Criteria for the Validation :- |  |
| 5.5.1 | - All validation activities shall be successfully performed and met with acceptance criteria defined in individual qualification protocols. - All discrepancies encountered during the execution of individual protocol shall be satisfactorily resolved. - The validation summary report shall be written and approved. |  |
| 5.6 | Re-validation Criteria :- |  |
| 5.6.1 | Re-validation of the Lab Solutions software shall be carried out by the external agency at an interval of every 5 years (± 1 month). |  |
| 5.6.2 | Complete / Partial re-qualification should be carried out under following circumstance with evaluation of change control request as per SOP No. QAGE09 (Change Control Procedure) :- - Change in server system. - Change in GxP Module of server. |  |
| 6.0 | ANNEXURES:- NA |  |
| 7.0 | REFERENCES:- NA |  |
| 8.0 | REVISION HISTORY:- |  |
| Sr. No.; Revision No.; Date of Revision; SOP No.; Supersedes No.; Reason for Revision. 01; 30/12/2022; QAGE133-01; QAGE133-0; - Change in Firm Name from “Indiana Ophthalmics” to “Indiana Ophthalmics LLP”. - Addition of Scope, Annexures, References and Revision History as Section 2.0, 6.0, 7.0 and 8.0 respectively in the SOP.. |