| 1.0 | OBJECTIVE : | To describe procedure for password policy, user rights information of HPLC system and audit trial review. |
| --- | --- | --- |
| 2.0 | ## SCOPE : | This procedure is applicable for password policy, user rights and audit trial review of HPLC system installed in QC department of Indiana Ophthalmics LLP. |
| 3.0 | ## RESPONSIBILITY : | QC Chemist, QC executive, QC sr. executive, QC Head / QA Designee, QA Head |
| 4.0 | ### ACCOUNTABILITY : | QA Head |
| 5.0 | PROCEDURE : |  |
| 5.1 | Each user shall have unique Username and Password for HPLC system. |  |
| 5.2 | In QC department, Lab solution software is used for operating HPLC system. |  |
| 5.3 | Password Policy for HPLC System :- |  |
| 5.3.1 | The Quality Assurance shall assign individual user & passwords to the different persons to access the system. |  |
| 5.3.2 | The user rights shall be categorized on the basis of access assigned to the given person. |  |
| 5.3.3 | The password shall be allocated at Administration level, Reviewer level and User level depending upon the availability of the provision in the HPLC system. |  |
| 5.3.4 | Password shall have at least 8 characters. |  |
| 5.3.5 | Password validity shall be of 90 days. |  |
| 5.3.6 | System shall not acquire last 3 expired passwords. |  |
| 5.3.7 | Account will get locked automatically after 3 wrong login attempts. Lockout of user shall be unlocked only by administrator. |  |
| 5.3.8 | The individual shall maintain the confidentiality of passwords assigned to them. |  |
| 5.3.9 | The password allocated shall be recorded in the Password Allocation Record as per Annexure-I and maintain it under lock and key in Quality Assurance department. |  |
| 5.3.10 | The access rights shall be defined by Quality Assurance department and shall be recorded in the Rights Allocation Record as per Annexure-II. |  |
| 5.4 | User Rights Password :- |  |
| 5.4.1 | Administrators: Responsible person denoted as Administrator shall be the member of this group. The members of this group shall have maximum rights and also shall have right to assign the privileges to the other privilege group. |  |
| 5.4.2 | Reviewer: Head QC / Head QA / Designee shall be the member of this group. The member of this group shall have rights such as create / edit user, create / edit method files, view audit trail etc. |  |
| 5.4.3 | Test Manager: The person responsible for the method preparation & generation of histogram shall be the member of this group. Test Manager shall have rights such as creation of batch table, set reference results, print reports, method create, create multi data report, edit report format, view audit trail etc. |  |
| 5.4.4 | User (Operator): The person responsible for the generation of histogram shall be the member of this group. User shall have rights such as creation of batch table, set reference results, print reports etc. |  |
| 5.4.5 | The privileges of the HPLC system (Lab solution software) to Administrator, Reviewer, Test Manager and User (Operator) shall be assigned as per Annexure-IV (HPLC (Lab Solution Software) Privileges). |  |
| 5.5 | Procedure for Audit Trail Review :- |  |
| 5.5.1 | The purpose of an audit trial is to make electronic data secure and traceable. Audit trial shows name of the person who assessed the system with date and time. |  |
| 5.5.2 | The HPLC audit trial shall include, the user name, date and time of the run, the integration parameters used and, when applicable, Invalid attempts to log on the system, further details on processing etc. |  |
| 5.5.3 | After completion of analysis QC reviewer shall review the data and shall review the audit trail of each analysis as part of review of analytical data as per applicability on daily basis. |  |
| 5.5.4 | Audit trial log shall be reviewed by the QC reviewer at every 07 ± 2 working days for any discrepancies related to data integrity and data security. |  |
| 5.5.5 | Reviewer shall login in to the Lab Solution software with their individual login ID and password. |  |
| 5.5.6 | Click on configure system and software window will open. |  |
| 5.5.7 | Click on administrator & and open to log browser. Audit trial display on. |  |
| 5.5.8 | Enter the date of analysis and user name and time period and event type. |  |
| 5.5.9 | Reviewer shall also check, for any abnormality observed related to audit trail deletion, renaming, alteration and movements, it shall be investigated and justified. |  |
| 5.5.10 | QC Reviewer review and check all activity done by analyst. |  |
| 5.5.11 | During review, QC reviewer shall analyze the difference between method versions to determine any significant change that has been made to the existing version. |  |
| 5.5.12 | QC reviewer shall also review any change in the system configuration, data archival history, any change in addition or deletion of user groups / systems. |  |
| 5.5.13 | Analysis date, time of analysis, total run injections with respect to system sequence print shall be reviewed during audit trial. |  |
| 5.5.14 | Log in and Log out timed events of users shall also be assessed during audit trial review. |  |
| 5.5.15 | If any abnormality or discrepancies are observed during the review, same shall be immediately inform to QC Head. |  |
| 5.5.16 | QC Head shall immediately inform to QA Head and the impact shall be evaluated for data integrity and data security. |  |
| 5.5.17 | After reviewing the audit trial log, QC reviewer shall complete the necessary entries in the “Audit Trail Log Review Record” as per Annexure-III. |  |
| 5.5.18 | QA Head / designee shall ensure the effective review of audit trial log as per defined frequency and shall sign in the respective column of “Audit Trail Log Review Record” as per Annexure-III. |  |
| 6.0 | ANNEXURES:- |  |
| 6.1 | Annexure-I : Password Allocation Record |  |
| 6.2 | Annexure-II : Rights Allocation Record |  |
| 6.3 | Annexure-III : Audit Trial Log Review Record |  |
| 6.4 | Annexure-IV : HPLC (Lab Solution Software) Privileges |  |
| 7.0 | REFERENCES:- NA |  |
| 8.0 | REVISION HISTORY:- |  |
| Sr. No.; Revision No.; Date of Revision; SOP No.; Supersedes No.; Reason for Revision. 03; 28/09/2022; QAGE77-03; QAGE77-2; - Change in Firm Name from “Indiana Ophthalmics” to “Indiana Ophthalmics LLP”. - Addition of Annexures, References and Revision History as Section 6.0, 7.0 and 8.0 respectively in the SOP.. 04; 02/10/2024; QAGE77-04; QAGE77-03; - Addition of Scope as Section 2.0 in SOP. - The word “Delete” removed from section 5.4.2 due to typo error. - Audit trial review period changed from 30 + 5 days to 07 ± 2 working days under section 5.5.4. - Annexure-IV added as “HPLC (Lab Solution Software) Privileges”.. 05; 17/11/2025; QAGE77-05; QAGE77-04; - Procedure for audit trial review elaborated in detailed. - QC reviewer added instead of QC Head in procedure. - Addition of section 5.5.3, 5.5.5 to 5.5.10, 5.5.13, 5.5.15 & 5.5.16. - QA Head / Designee added under section 5.5.18 instead of QA Head. - Annexure-III updated with reviewed by QC reviewer / designee instead of QC Head and designee added in verified by QA head column.. |